22.058, Principles of Medical Imaging

Fall 2002

Homework #4

Due Tuesday October 22nd

________________________________________________________________________

1. Write a complete system description for the instrument function of a planar x-ray imager (assume scanned fan beam).  Include:

• finite size source

• heal effect on source intensity and energy spectrum

• oblique angle effects

• depth dependent magnification

• quantum efficiency and PSF for the scintillator/photographic plate.

- To include energy effects, make:

  - is a function of E

  - add integration of E to both integrals

- Quantum efficiency degrades I by, a uniform effect

2. For a cylindrical object (long axis perpendicular to the beam) calculate the profile of X-ray intensity in a fan beam geometry, assuming that the beam is mono-energetic.

3. Calculate the effect of beam hardening on the CT image of a disk.

The center is less attenuating than it should be, therefore the image is:

4. For the following sample, show (a.) the projections and (b.) the filtered projections.

See Appendix A.

5. A sinusoidally modulated x-ray image is recorded by a one-sided screen film system as shown below.  Find the recorded S/N as a function of frequency, where the signal is the sinusoidal component and the noise is the average background.  On average the screen produces l photons per x-ray photon, t of which are transmitted to the emulsion where r are recorded.  The pixel area of the film is much smaller than the system resolution.  Neglect any critical angle effect between the screen and the film.

x-ray photon number as a function of z = n0 (1+cos(2 π k z)).

6. Write a program that calculates the Radon transform of an object function, then Fourier filters the projects, and finally reconstructs an image via back projection.

See Appendix A.

# APPENDIX A: Mathematica File (Projection2.nb)

## Projection reconstruction and the Radon Transform 2

### The Radon Transform

The forward Radon transform is to convert a 2-D object into a set of projects within the plane.

Radon[object_, n_, fov_] :=

Table

[

Integrate

[

object DiracDelta

[

m - x Cos[\[Theta]] - y Sin[\[Theta]]

],

{x, -fov, fov},

{y, -fov, fov}

],

{m, -2 fov/(n - 1),

+2 fov/(n - 1), 2 fov/n},

{\[Theta], 0, Pi, Pi/(n - 1)}

];

The double integral on the previous page is very slow to evaluate, and so we reduce it to a line integral along the line defined by the delta function.

Radon2[object_, x_, y_, n_, fov_] :=

Table

[

Nintegrate

[

object[x, y],

{yp, -fov, fov},

{PrecisionGoal -> 4}

],

{xp, -fov, fov, 2*fov/(n - 1)},

{\[Theta], 0, Pi, Pi/(n - 1)}

] // N

### Define a simple test object

object1[x_, y_] := If[x^2 + y^2 < 256, 1, 10^(-6)] // N;

Plot3D

[

object1[x, y],

{x, -64, 64},

{y, -64, 64},

{PlotRange -> All, PlotPoints -> {64, 64}}

]

Robject1 =

Radon2

[

object1,

xp Cos[\[Theta]] - yp Sin[\[Theta]],

yp Cos[\[Theta]] + xp Sin[\[Theta]], 64, 64

];

ListPlot3D[Robject1]

### Filtered Back Projection

Fdata = Fourier[Robject1];

ListPlot3D[Re[Fdata], {PlotRange -> All}]

Filt =

Table

[

If

[

x <= 32 && y <= 32,

Sqrt[x^2 + y^2],

If

[

x <= 32 && y > 32,

Sqrt[x^2 + (65 - y)^2],

If

[

x > 32 && y <= 32,

Sqrt[(65 - x)^2 + y^2],

If

[

x > 32 && y > 32,

Sqrt[(65 - x)^2 + (65 - y)^2]

]

]

]

],

{x, 0, 63},

{y, 0, 63}

];

ListPlot3D[Filt]

FiltFdata = Fdata*Filt;

ListPlot3D[Re[FiltFdata], {PlotRange -> All}]

Filtdata = Fourier[FiltFdata];

ListPlot3D[Re[Filtdata], {PlotRange -> All}]

### Back Projection of Filtered

Bflimited[x_, y_, n_] :=

If

[

x^2 + y^2 > 32^2,

0,

1/(2 Pi)

Sum

[

Transpose

[

Re[Filtdata]

]

[[m*64 + 1]]

[[Floor[x Cos[m*Pi] + y Sin[m Pi]] + 33]],

{m, 0, 1 - 1/n, 1/n}

]

];

Plot3D

[

BFlimited[x, y, 4],

{x, -32, 32},

{y, -32, 32},

{PlotRange -> All, PlotPoints -> {64, 64}}

]

DensityPlot

[

BFlimited[x, y, 4],

{x, -32, 32},

{y, -32, 32},

{PlotRange -> All, PlotPoints -> {64, 64}, Mesh -> False}

]

Image =

Table

[

BFlimited[x, y, 64],

{x, -32, 32},

{y, -32, 32}

];

ListPlot3D[Image, {PlotRange -> All}]

ListDensityPlot[Image, {PlotRange -> {0, 500}, Mesh -> False}]