How to Run an MFP Experiment

I.  Initialization

1.  Center the MFP on the vibration table.  Turn on Vibration table by flipping

the switch in the center of the back of the table.  Use the hexagonal wrench 	to adjust the two screws on each side such that all four pins are in the center 	of their holes.

2.  Turn on the laser by flipping the switch on the MFP controller.

3.  Create a file and backup folders for the experiment you are about to run.

4.  Open a new MFP experiment by double clicking on one of the packaged box

icons on the computer screen.  Go to “MFP Controls” and open up “Pulling Controls”, “Sum and Deflection Meter”, “Start Hampster”, and “Video”.

5.  In the Pulling Controls box, click on “Save to Memory” and “Save to Disk”.  	For the latter, browse to select the relevant backup folder you created in step 	3.

- Enlarge the video screen for better viewing.

II.  Focusing the laser on the back of the cantilever

1.  Clean the tip holder either by rinsing and wiping with the solvent you are about 			to use, or with ethanol if you are not using a solvent.  The plastic part is of 			importance here, not necessarily the metal.  Alternatively, you may 				ultrasonicate the tip holder for 10 minutes to get a very deep cleaning.  This 			may be done by filling the sonicator with water to the fill line, placing the tip 			holder in a small beaker and cover it completely in the appropriate solvent 			(the one you will be using or ethanol), then placing the beaker in the 				sonicator bath and turning the dial to 10 min.

2.  Break off a new chip by first placing a row of chips, tip-side up, onto a glass 			cover slip which has been fixed to a glass slide by tape, e.g.  Position the row 			such that only one chip is extending beyond the edge of the cover slip 			leading to the far end of the glass slide.  Holding down the second chip from 			the end extending over the edge with a pair of tweezers, apply pressure to the 		center of the backside of the chip over the edge until it break off the row and 		lands safely on the glass slide.  Place the rest of the row back into the gel 			pack.  Remember that the tips are on the upside of the cantilever (the side 			with the Omega sign) So If a chip ever flips over, most likely you have ruined the tips 			or at the very least contaminated them.  Throw the chip away and begin again!

3.  Place the new chip into the tip holder by pressing gently back onto the spring-			loaded gold tip clamp, sliding the chip underneath it (tip-free end first), 			being careful not to touch the tips, and then release the tip clamp.  Nudge 			the chip gently on either side to ensure it is secure.  Invert the MFP head, 			load the tip holder, and replace the head to the microscope base.

4.  Shine the snake light at the chip from the side so that it is illuminated on the 			video screen.  Focus and position the microscope so you can clearly see the 			tip you are about to use.  Make sure it doesn’t look broken or contaminated!  			Now move the light slightly away so you can see the laser beam on the 			screen.  Focus the laser on the back of the cantilever by first positioning 			the laser on the edge of the chip between the legs of the cantilever, then 			moving it out onto the tip of the V.  Tweak the laser positioning on the tip 			such that the “Sum” in the Sum and Deflection Meter is maximized.  			Finish by turning the Photo Diode knob until the “Deflection” in the Sum 		and Deflection Meter reads zero.  Move the light away and turn it off.  			Close the video window.

III.  Calibrating the Tip

A.  If you will be working in fluids:

1.  Invert the MFP head and remove the tip holder.  Wet the tips with a 				small drop of the solution you will first be using for your experiment.  			Verify that the tips are	wet by viewing them with the monocular.  				Replace the tip holder in the MFP head.  Do not yet re-invert the 				MFP head.

2.  Fix a piece of mica to a glass slide using vacuum grease or tape, then 				cleave off the top sheet of the mica with a piece of tape.  Wet the 				mica with 2-3 drops of the fluid you will be using and place it on the 				sample stage.

3.  Being careful not to move the laser adjustment knobs, gently 				replace the MFP head, back feet first, onto the base, and slowly 				lower the front foot while watching the tip holder and sample to 				ensure that the tip does not crash into the sample.  Adjust the front 				foot if necessary.

4.  Lower the front foot until a meniscus forms with the fluid, then gently 				re-adjust the laser positioning to re-maximize the sum and re-zero 				the deflection.

5.  Lower the front foot until the tip is just above the surface of the 				mica.  Do this by lowering until the deflection meter just turns red, 				then backing off until the tip is just free from the surface (the 					deflection meter should read blue or zero.  Be careful 	to ensure the 				tip is really free, and not sticking.  Re-zero the deflection meter using 				the photodiode knob on the MFP head.

6.  In the “MFP Display” menu, select “Graph Options”.  For the y-axis 				select “Deflection” and for the x-axis select “LVDT”.  Select 					“Realtime Graph”.

7.   Make sure the little black switch on the MFP controller is switched to 				“MFP Z” then take a force curve by clicking “Do Single Pull” on 				the pulling controls menu.  If the resultant curve is all noise, the tip is 			completely off the surface.  If the resultant curve is thin and linear 				with a negative slope, the tip is completely on the surface.  Turn the 				hamster a full turn or so to lower the piezo toward the surface, then 				try the force curve again.  Repeatedly adjust the hamster and taking 				force curves until you get a curve that has a non-contact regime 				and a constant-compliance regime.

8.  Determine the correct IOLS Value for the tip:  Create a review graph 			by going to the “MFP Display” menu and selecting “Make Review 				Graph”.  In the review panel to the left of the 	Review Graph, select 				the suffix of the graph you wish to analyze from step 7.  Note that 				the suffix of the next curve to be taken is displayed in the control 				panel under “Scan Controls”, so the suffix of the last curve taken is 				one less than this number.  Select the Review Graph and type 				<ctrl><i>:  two cursors, one round and the other square, should 				appear in the lower left hand corner of the Review Graph.  Click and 				drag each of them such that they are ~2cm apart on the blue 				retraction curve, and in the constant compliance regime near 				the point of contact.  Go to the “Analysis” menu and select “Curve 				Fitting”.  In the “Function and Data” menu under Curve Fitting, 				select “line” for Function, and click on “From Target”.  Then for the 			“Y data” select the deflection curve you are analyzing, and for the “X 			data” select the LVDT curve you are analyzing.  Then go to the 				“Data Options” menu in the same box and select the “Cursors” 				button.  Finally, click “Do It” at the bottom.  You have just 					calculated the slope of the line.  In the dialog box there will be a line 				four or five lines from the bottom beginning with “W_coef”.  You 				want the second number in the brackets on this line.  Divide the 				current IOLS value (located at the bottom of the pulling controls 				box) by this number.  Igor can do this for you if you type in the 				dialogue box <print XXX/YYY><enter>, where XXX is the 				current IOLS value and YYY is the second number from the W_coef 			line.  The result of this calculation should be within the range of 60-				150nm/V.  If it is not, there is most likely something wrong with the 				tip.  Enter the result into the IOLS window in the pulling controls 				menu.  Igor will save this new entry until it is changed by another 				user or yourself.

9.  Calibrate the spring constant of the tip:  Ensure the tip is not too 				close to the surface (i.e. no interaction forces are detected), as this 				will damp out the thermal spectrum, then switch the little black 				switch on the controller to “User Z”, and click on “Thermal PSD” 				in the pulling controls box.  Wait for it to complete the process, 				about 20 seconds.  Then go to the “Graph” menu and select 					“Autoscale Axes”  then double click on each axis and select “Log” 				and then “Do It” in the box that pops up for each.  Now enter the 				theoretical resonance frequency into the Resonant Frequency 				window in the left hand column of the Thermal Panel box, and click 				“Fit Me”.  (Note that the theoretical 	values for the 						Thermomicroscopes tips are printed on a sheet on the MFP 					corkboard on the wall of the lab.  Be sure you are reading the 	correct 				value for the tip and conditions - air or water - you are using.  Also 				note that the values on this sheet are in KHz, while the Thermal 				Panel box is using Hz.) The red line in the thermal spectrum should 				fit itself to the true primary resonance frequency of the tip, and the 				true spring constant will appear both in the Thermal PSD box and in 				the Pulling Controls box.  The more times you repeat this process of 				entering the theoretical	resonance frequency and clicking “Fit me”, 				the better the fit will be and soon you will converge onto the most 				accurate value for the spring constant.  The spring constant should 				not be more than ±½ of the theoretical value.  If it is, most likely 				there is something wrong with the tip.  Print out the thermal 				spectrum and make note of the spring constant, resonance 				frequency, Inverse Optical Lever Sensitivity (IOLS), and the 				Sum.  Save the Calibration file and close Igor.

B. If you are working in air:  Skip steps 1 and 4 from the above instructions for 			working in fluids, and for step 2 you of course will not wet the mica.  All 			other instructions are the same.

IV.  Running an Experiment

1.  Create a backup folder for your experiment if you have not already done so.

2.  Open a new MFP experiment by double clicking on one of the packaged box

icons on the computer screen.  Go to MFP Controls and open up 				Pulling Controls, Sum and Deflection Meter, and Start Hampster.  You do 		not need to open up the Video this time, since the laser is already focused on 		the tip.

3.  In the Pulling Controls box, click on “Save to Memory” and “Save to Disk”.  	For the latter browse to select the relevant backup folder you created in step	3.

4.  Most samples may be fixed onto a glass slide with tape or vacuum grease.

5.  Wet sample if necessary with the appropriate fluid.

6.  Lift up the front foot of the MFP head, replace the mica slide with your sample, 			and gently lower the front foot again, careful to not crash the tip.  Be warned 		that most samples are thicker than mica and the front foot will thus usually 			have to be raised considerably for this step!

7.  Bring the tip down close the surface as in step III.A.5.

8.  Adjust parameters such as Pulling Distance, Scan Rate, and Points per Second.  			For single molecule measurements, I typically use a Pulling Distance of 			1.0µm, a Scan Rate of 1.0µm/s, and Points per Second data collection of 			5000.  Your own parameters may vary.

9.  Go to the MFP Display menu and select the “Graph Options” menu and select 			“Force” for the Y-axis and “Z-Position” for the x-axis. Select Realtime 			Graph.

10.  Switch the little black switch on the contoller back to “MFP Z”.
	11.  Engage the surface and begin taking for curves!  You may review the curves you 			have taken at any time with the “Make Review Graph” option under the 			“Graph Options menu.

V.  At the end of the experiment

1.  Save your experiment!  Close Igor.

2.  Remove and properly store your sample.

3.  Remove and properly store your chip, as the tips are often needed for subsequent 			characterization and analysis.
	4.  Clean the tip holder with the appropriate solvent.

5.  Turn off the controller and disable the vibration table.

6.  Clean the workspace so that it is as clean or better than it was when you arrived.

7.  Be sure to sign the logbook and note any difficulties!

Written on 19 December, 2001 by Monica Rixman