## Slide 1: FSC-A Scaling

## Slide 2: Area and Height of FSC pulses must be the same.
45°

## Slide 3: WHY?

## Slide 4: The FSC threshold value is based on the FSC pulse height. NOT area
The value of 5000 is not relevant on the FSC-A axis since FSC-A ≠ FSC-H by default

## Slide 5: WHY does FSC-A≠ FSC-H?
The laser is only 8 um high and the cell is larger than the 8 um
EXCEPTION, these cells are not larger than 8um: bacteria, yeast and RBC’s or platelets

## Slide 6
The pulse flattens out on top because the cell is larger than the laser.
Without being fully illuminated the FSC-H pulse can never reach its true maximum
FSC-A will always be equal to the full FSC value of the cell, therefore FSC-A> FSC-H IF the cell is larger than the laser
Chart: time
 | 0.2 | 0.5 | 1 | 4 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 10 | 4 | 3 | 2 | 1 | 0.1 | 0
B
C
D
A
B
C
D
laser

## Slide 7: You must multiply the FSC-A by some number less than 1 to make the FSC-A = FSC-H

## Slide 8: HOW?

## Slide 9: Adjust the FSC-A Scaling number so the cells fall on a 45° angle on a FSC-A vs. FSC-H plot

## Slide 10
BEFORE
FSC Area Scaling adjustment
FSC Area Scaling= 0.75
AFTER
FSC Area Scaling adjustment
FSC Area Scaling= 0.42

## Slide 11: Now the threshold value of 5000 has relevance on a FSC-A vs. SSC-A plot
= debris
P1= live cells
= dead cells

## Slide 12: In conclusion…
FSC threshold is based on the FSC-H
Adjust the FSC-A using the FSC-A Scaling under cytometer window> laser tab
So FSC-A= FSC-H by placing the cells on a 45° angle on a FSC-A vs. FSC-H plot